## Slide 1: Transcatheter intervention in extremely low birth weight (ELBW) infants
Mohamad-Taghi Majnoon, MD
Pediatric Cardiologist
Assistant Professor
Children’s Medical Center of Excellence
Tehran University of Medical Sciences

## Slide 2
PDA is common in preterm neonates and can destabilize cardiorespiratory status.
Traditional options: medical therapy or surgical ligation.
Transcatheter device closure offers a minimally invasive alternative.
Piccolo enables percutaneous closure in ELBW infants ≥700 g.
Background & Scope
1/19

## Slide 3
Can worsen RDS, prolong ventilation, and drive cardiovascular instability.
Associated with higher risks of BPD, IVH, renal dysfunction, and mortality.
Clinical impact greatest in infants <28 weeks’ gestation.
Yet ~85% of untreated PDAs close spontaneously after discharge.
Hemodynamically Significant PDA (hsPDA):
Why It Matters
2/19

## Slide 4
No single consensus on echocardiographic or clinical thresholds.
Marked variability across institutions and countries.
Pulmonary vascular resistance evolves over days–weeks postnatally.
Requires serial reassessment of hemodynamic significance.
Defining hsPDA: Ongoing Controversy
3/19

## Slide 5
Combine echo assessment of shunt size/volume with clinical consequences.
Evaluate ductal dimensions and morphology within overall hemodynamics.
Integrate ventilatory needs, perfusion markers, and end‑organ effects.
Standard approach remains multi‑parameter and iterative.
Assessing Significance
4/19

## Slide 6
1967 Porstmann: first transcatheter PDA closure.
Long limited by access/device constraints in tiny infants.
ADO II expanded feasibility in larger infants; later devices pushed limits.
Closure now possible in ~1–2 kg infants at experienced centers.
Evolution of Transcatheter PDA Closure
5/19

## Slide 7
FDA‑approved (2019) for patients ≥700 g; engineered for ELBW anatomy.
Piccolo used in 54.8% of procedures;
Prospective multicenter (≥700 g) shows excellent 3‑yr safety/efficacy.
Amplatzer Piccolo Occluder—What Changed?
6/19

## Slide 8
Weight: ≥700 g (Piccolo); often ≥1500 g for other devices.
Postnatal age: typically ≥3 days for stabilization.
Echo‑confirmed L→R shunt with hemodynamic consequences.
Failed medical therapy (indomethacin/ibuprofen/acetaminophen).
Anatomic suitability (e.g., duct 2–5.4 mm) and clinical stability.
Patient Selection—Typical Inclusion
7/19

## Slide 9
Standard: femoral venous access; fluoroscopic + echocardiographic guidance.
Bedside occlusion avoids transport risks in fragile ELBW infants.
Reports of faster recovery and reduced transfusion needs at bedside.
Team expertise and infrastructure determine feasibility.
Cath Lab vs NICU Bedside Closure
9/19

## Slide 10
Define ductal anatomy and select device size/type.
Advance device via venous route across ductus under echo/fluoro.
Deploy with attention to LPA and aortic origins to avoid obstruction.
Confirm stability and residual shunt; finalize release.
Technical Steps (Typical Workflow)
10/19

## Slide 11
Device embolization (~2.7%); often amenable to percutaneous retrieval.
Left pulmonary artery (LPA) stenosis; may require stent.
Aortic obstruction from device protrusion (uncommon with proper sizing).
Tricuspid regurgitation from device–valve interaction.
Complications—Major Device‑Related
11/19

## Slide 12
Hemolysis (rare), infective endocarditis (rare), residual shunt.
Consensus guidance published for prevention and management.
Emphasize meticulous technique and vigilant post‑procedure monitoring.
Other Complications & Risk Factors
12/19

## Slide 13
Match device to ductal size/morphology; avoid protrusion.
Continuous echo guidance to surveil LPA/aortic patency.
Plan for retrieval strategies if embolization occurs.
Structured post‑closure surveillance protocols in NICU/clinic.
Preventing & Managing Complications
13/19

## Slide 14
Transcatheter: rapid recovery, fewer transfusions, no incision; bedside possible.
Surgical: decades of outcomes; useful with anatomic contraindications; accessible.
Utilization: >80% invasive closures are catheter‑based in some centers.
Further high‑volume comparative data still needed.
Transcatheter vs Surgical Ligation
14/19

## Slide 15
Large series: 758/762 (99.5%) successful implantation across devices.
Device mix: ADO II 38.8%, Piccolo 54.8%, AVP II 5.8%.
Complete ductal closure common on follow‑up
Durable safety/efficacy reported in multicenter Piccolo cohort.
Clinical Outcomes
15/19

## Slide 16
~23% had LVEF <55% at 24 h; all normalized before discharge.
Pre‑procedure LV dilation predicts transient LVEF reduction.
Plan monitoring and support in early post‑closure period.
Left Ventricular Function After Closure
16/19

## Slide 17
Study (Lai et al.): no significant BPD benefit vs medical management.
Possibilities: closure not early enough or population selection effects.
Optimal timing remains unsettled; early targeted therapy for select high‑risk infants may be considered.
Pulmonary Outcomes & Timing
17/19

## Slide 18
Agents: indomethacin, ibuprofen, acetaminophen remain first‑line.
Many persist despite prophylaxis; some need further therapy/ligation.
Many PDAs close spontaneously; trend toward conservative initial management.
Transcatheter/surgical closure as rescue for refractory hsPDA.
Medical Therapy & Contemporary Practice
18/19

## Slide 19
Meticulous anesthesia/airway and ventilation strategies are essential.
For interhospital transport: stabilize, monitor, and coordinate teams.
Bedside transcatheter closure reduces transport‑related risk when available.
Piccolo enables closure ≥700 g with high success and low complications in experienced hands.
Outstanding questions: optimal timing, neurodevelopmental impact, cost‑effectiveness.
Anesthesia, Transport & Key Takeaways
19/19

## Slide 22: Thanks for Attention